LAB:  FERRIS 213

Lab 1, Introduction to the Spectrum Analyzer (SA)

ECE 341 Spring 2010

Objectives:

- Become familiar with functions of the lab R.F. signal generator (RFSG).

- Become familiar with functions of a modern spectrum analyzer (SA).

- Measure some actual characteristics of a sine wave.

- Measure the spectra of FM radio signals and (digital) TV using the SA.

Some background:

- Power calculation:

- , where V is RMS voltage and R is resistance. In many systems, R is 50 ohms.

- Decibel representation (log of the ratio of the quantity to be represented, to another quantity):  which is the same as .

- One representation adds “m” to mean “dB miliwatts”: .

- One can also use dB(A to compare a signal to microamperes, dBmV to compare a signal to millivolts, etc.

- On the SA, the frequency (horizontal) axis is specified by either the center frequency and span, or by the start and stop frequencies.

- On the SA, the vertical axis can be linear or log (the default). The reference level is the value at the top of the screen, and the scale is given by the value per division (default: “10 dB/” for 10 dB per division.

- On the SA,  “BW/Avg>Res BW” lets you set the resolution bandwith, which is simply the width of the filter scanning the spectrum of the signal.

- Also “BW/Avg>Res BW” provides viewing the average of the last N scans, which usually helps signal stand out from background noise. You specify N.

- The main buttons we use on the SA are FREQUENCY, SPAN, AMPLITUDE, BW/AVG, SAVE, and MARKER. Numbers can be entered by the number pad or by the wheel. Once a number is put in, units can be chosen (right side of screen).

Precautions:

- Never tighten connectors very tight!

- The SA input is very sensitive and easily damaged. To be sure the cable has no charge before connecting it to the SA, connect the cable first to the source (to discharge the cable) and then to the SA.

Start Up:

- Turn on the RFSG and SA so that they warm up.

- Be sure the RFSG has its RF set to OFF to begin, and while adjusting the output. Otherwise the wrong RFSG output could damage the SA.

- Connect the RF OUTPUT port of the RFSG to the INPUT 50( port of SA with the 2ft, 50( Agilent cable.

Part A: sinusoidal signal

(Each step is an action for you to take. Bold words in capitals are usually instrument buttons. The “>” is like “then; e.g. BW/Avg>RES BW>1kHz means the BW/Avg button, then the RES BW button, then put in 1 kHz.)

- RFSG settings:

  - Amplitude:  -10 dBm (See “Some Background” above for meaning of dBm.)

  - Frequency: 1250 kHz (or 1.25 MHz)

  - RF: ON. (While settings were being made, the output was not actually on.)

- SA settings:

- PRESET: Notice that the initial screen horizontal (frequency) axis has center frequency 1.5 GHz and Span 3 GHz, making the horizontal axis go from 0 to 3 GHz. Can you observe the desired 1250 KHz signal on the screen? Where should it appear on the screen with these settings?

o 	FREQUENCY>CENTER FREQUENCY> 1250 KHz.     Notice the peak is at 1250 kHz!

o 	BW/Avg>RES BW> 1 kHz. Wait for one scan to finish. Why is the scan slow? (Hint: Res BW = 1 kHz means the scan is done in steps of 1 kHz, over the entire span.)  Why is the spectrum of the 1250 kHz much narrower than before? Which is more desirable? Which is correct?

- SPAN>50kHz This spreads out the spectrum and we can see how “wide” it is.

- Position the signal peak so it and most of the entire curve can be seen well, as follows:

AMPLITUDE>REF LEVEL>select with knob  (Reference Level is the value of

the spectrum at the top of the screen.) Bring the peak of the spectrum almost to the

top of the screen so we can measure how wide the spectrum is.

  - Define markers 1, 2, and 3 by repeated use of (Marker>SELECT MARKER>NORMAL). Select marker 1 and use the wheel to position it on the spectrum peak. Select marker 2 and put it 50 dB lower then the peak, on the left of the peak. Select the 3rd marker and put it 50 dB lower then the peak, on the right of the peak. We can call the difference in the two extreme frequencies the “50 dB bandwidth” of the spectrum we see. What is the value? Is it reasonable for a good sine wave at 1250 kHz? (Ideally, of course, it should be zero, but there is always some width in spectra of real sine waves.)

  - Marker>More>Marker Table on   to get all data and graph on the screen. Check your frequencies and dB values.

  - Save the graph and data on the screen to a floppy disk. You can then use the lab computer to email the graph image to your group members.

    - Title the graph image file by the last names of your group members, separated by a comma: DISPLAY>TITLE>CHANGE TITLE> name1,name2.

    - Save the screen to the disk: FILE>SAVE>TYPE>SCREEN,   FORMAT>BITMAP,   NAME>LAB1AW,   RETURN,   DIR UP,   SELECT A DRIVE,   DIR SELECT,   SAVE NOW.

- Now we check another characteristic of the sine wave’s spectrum. START FREQ> 1 MHz,  STOP FREQ 4 MHz. Looking at more of the spectrum, we see there are several smaller peaks at quite different frequencies from the frequency we set on the RFSG! Use markers to measure the frequencies of all peaks. (If the RFSG output is not a pure sine wave it is still periodic, and then in ECE 315-316 we learn that we have frequencies at the desired one and integer multiples of it. What frequencies do you have?

- Get the marker table and save it to the disk as before; call the file LAB1AF.

PART B: FM RADIO SIGNALS

- Connect a suitable antenna to the 50 ohm input port of the SA with the 2 foot, 50 ohm Agilent cable. A minimal antenna is a short rigid wire on the ungrounded terminal of an adapter for BNC to banana jacks.

- Set the SA to display 87.0 MHz to 108.5 MHz to show the full FM radio spectrum,

- Place a marker at the peak of the strongest signal and read the value of the frequency shown on the screen. What is the frequency?

- Find a second station, perhaps your favorite or a UTK station, and similarly measure its center frequency.

- Save your graph to a floppy disk like you did in Part A, and email it to yourselves for your report. Title it LAB1B.

PART C: TELEVISION SIGNALS

- The lab instructor will specify the frequency of a channel for your group for your group.

- Connect a suitable antenna to the SA input. The lab TA can make suggestions.

- Obtain a display of the spectrum.

- Place a marker at each end of the spectrum and see what range it covers.

- Save the screen image and email it to yourselves as above. Title it LAB1C.

REPORT:

Each person should submit a separate report no more than a week after doing the lab. The lab TA will explain the format and how to submit the report.

The end.